## Slide 1: Smartcheck Functional Test Kit

## Slide 2: Introduction to SmartCheck Functional Test
Purpose of weight
  Measure the ability to generate similar results
  Ensure repeatable across multiple measures
  Eliminate all known systematic sources of error
Recommended frequency
  At least once per month, per Mettler Toledo/LabTech Weights and Measures recommendations
  May vary depending on lab guidelines, SOPs, and risk assessment such transport of unit

## Slide 3: SmartCheck Functional Test Kit SmartCheck Functional Test Kit includes all components needed to perform a SmartCheck functional test
Cover
1g weight
Weighing pan
Cleaning cloth
Forceps (tweezers)
Storage Case

## Slide 4: Perform a SmartCheck Functional Test Follow the steps shown to verify your SmartCheck functions properly
Power on SmartCheck then remove draft shield
1
2
Using tweezers, place weighing pan from your functional test kit atop vessel then cover
3
When 1000uL light blinks, remove cover, place the 1g weight on the pan then cover
4
Wait for 1000uL light to stop blinking

## Slide 5: Perform a SmartCheck Functional Test (continued) Follow the steps shown to verify your SmartCheck functions properly
Using tweezers, remove 1g weight then replace cover
5
7
Wait for 1000uL light to stop blinking then repeat – 10 measurements in total
Watch YouTube video: "How to check your SmartCheck"
8
Green lights indicate PASS, SmartCheck is ready for use
6
When 1000uL light blinks, remove cover, place the 1g weight on the pan then cover
Place black cover over 1g weight